## Slide 1
CBPM System Overview
(Part I?)

## Slide 2: Locations Of Key Components
All based in: $CESR_ONLINE (/nfs/cesr/online/CBPM)
  /instr
    /allocation
    /config
    /ops
    /log/CBPM
  /acc_control
    /bin

## Slide 3: Locations Of Key Components
$CESR_ONLINE/instr/allocation
  Instrument_allocations
    Named collections of instruments for operations and testing
    Holds allocations for West, East, and current monitor collections of instruments

## Slide 4: Locations Of Key Components
$CESR_ONLINE/instr/config
  BPM_DET_params.cfg
    Beam pipe dimension coefficients
    Turn offsets for arbitrary bunch phase measurements
    MPM data publishing addresses
  BPM_INST_params.cfg
    Locationhostname mapping
    All timing parameters
    Gain tables
    Pedestal storage
  Config file backups

## Slide 5: Locations Of Key Components
$CESR_ONLINE/instr/ops
  cbic.conf
    Path definitions for:
      Instrument software images, parameter files, data, allocation file, bunch pattern definition file
    Default gain setting
    Default timing setup
  CBPM_II.ldr
  gain_cal_mapping.in
    Used to load gain calibrations produced by BPM_tbt_gain software
  Instrument software backup directory

## Slide 6: Locations Of Key Components
$CESR_ONLINE/instr/log/CBPM/<year>
  Log files that archive all terminal output from servers started via the start_WEST[EAST] method.
  Useful assigning a time to certain problem occurrances

## Slide 7: Locations Of Key Components
$CESR_ONLINE/acc_control/bin
  cbic
    symlink to server executable in CBIC archive directory
  /CBIC
    directory containing version history of executables
  start_BPM_server
    Wrapper script for starting the default server configuration for East and West
  start_EAST start_WEST
    Symlink to wrapper script. Runs East/West server.
  start_CURRMON
    Brings the 4ns bunch-by-bunch current monitor online

## Slide 8
Start Up

## Slide 9: Procedures
Starting the servers
  On a CESR online machine
    Typically terminal sessions are started on cesr105 (West) and cesr106 (East). Documentation for this is posted next to BPM station in the control room South.
  Server programs have three modes
    Interactive (menu) mode
    Server mode
    Current monitor mode
  Default mode when starting via start_WEST[EAST] is server mode.

## Slide 10
Server Mode

## Slide 11: Server Mode
Waits for measurement requests to be presented via MPM mailbox array
  A protocol for requesting and honoring measurements is implemented in the mostly unused CesrBPM library.
    Proper maintenance habits dictate that this measurement request protocol and implementation should be migrated into an existing CBPM II system library.
    The CesrBPM library pulls in mountains of obsolete machinery and memory in addition to the simple measurement request mechanism.
    This is kept around for measurement request support on VMS. Should be done away with ASAP.

## Slide 12: Server Mode
Measurement types available
  “orbit”
    Provides average button values and positions calculated via geometric coefficients found in BPM_DET_params.cfg
    Publishes values to MPM locations
    Programs known to request this measurement
      cesrv
      orbmon (a packaged cesrv)
      “orbit”?

## Slide 13: Server Mode
Measurement types available
  “turn-by-turn (TBT)”
    Produces one “RD” data file per active server
    Programs known to request this measurement
      cbpm_tbt_positrons / cbpm_tbt_electrons (VMS only)
        Actually a DCL wrapper script around a program built using the CesrBPM library on VMS.
        No idea where the source to this program lives
    NOTE: Active servers will honor the last predefined bunch pattern installed via the command 38) Install Turn-By-Turn Bunch Pattern
    Depending on the timing setup, this may control species.
    A new implementation and protocol for use is needed as this is prone to user error.

## Slide 14: Server Mode
Measurement types available
  Betatron phase measurement
    Publishes phase-related and amplitude-related values to MPM
    Programs known to request this measurement
      cesrv

## Slide 15: Server Mode
Measurement types available
  Pedestal acquisition
    Acquires new pedestal values for all timing setups supported by the instrument.
    Does not save to disk
    Programs known to request this measurement
      none

## Slide 16: Server Mode
Measurement types available
  Time scan (auto time)
    Acquires updated delay paramters for the currently active timing setup
    Does not save to disk
    Programs known to request this measurement
      cbpm_adjust_times (VMS only)

## Slide 17
Interactive Mode

## Slide 18: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  3) test data buffers
    Useful for determining if the instrument’s DSP is operating correctly
  4) get turn by turn data
    A manually requested TBT acquisition with more flexibility in requested parameters than server mode provides
  9) save timing parameters
    Writes all locally cached timing parameters to a temporary file for possible review
  10) save pedestal parameters
    Does the same for locally cached pedestal values

## Slide 19: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  11) time scan
    Prompts for several options allowing adjustment of delay values to optimize signal sampling
    Allows for plotting of waveforms found in a given bunch slot
  14) reread parameter files
    Replace cached delay, pedestal, and gain parameters with those from disk and push to all active instruments
  15) select active instruments
    Allow selection of specific instruments to obey future commands

## Slide 20: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  17) enter server mode
    Does what it says on the tin
    <CTRL><C> breaks out of server mode back to menu
  19) acquire pedestals
    Intended to be run with no current in CESR
    Acquires no-signal ADC value averages for all timing setups
  20) dump raw data
    Performs all actions normally taken in a TBT measurement, except acquisition of new values

## Slide 21: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  22) set timing setup
    Changes the timing setup state of all instruments currently active
    Causes delay values to be loaded into operational registers on active instruments
  23) install updated parameters
    Copy temporary parameter file created by save timing/save pedestal commands to master parameter file.
  25) program flash
    Installs the instrument code image specified in the application config file into all active instruments

## Slide 22: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  26) reset and init
    Restart the instrument software via a DSP reset
    Push all configuration information to instrument
    Set timing setup to last known value
  27) power cycle – reset and init
    Perform a power cycle of individual instruements
    Then perform the same actions as reset and init
  28) time-in
    Re-acquire all timing parameters from scratch on bunch 1 of the target species
    Prompts for the timing setup to use

## Slide 23: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  30) apply gain mapping corrections
    Read tables provided in the gain_cal_mapping.in file and apply them according to the method specified in the header of that file. (absolute or multiplicative)
    gain_cal_mapping.in’s contents must be copied from the output directory of BPM_tbt_gain software after its use.
  31) compose bunch pattern
    Displays a GUI for composing and saving to disk a custom bunch pattern for use in any timing setup

## Slide 24: Interactive (Menu) Mode
Calibration, and maintenance, and testing commands
  36) exit
    Terminate program
  38) install turn-by-turn bunch pattern
    Prompt for the saved bunch pattern to honor for all future TBT measurements

## Slide 25
Current Monitor Mode

## Slide 26: Current Monitor Mode
Start up
  Current monitor is the same program, but is started with different options.
    start_CURRMON handles this
    Current monitor will start in continuous acquisition mode and will publish signal values for ~620 bunches to MPM
Exiting
  One can exit continuous monitoring mode with <ctrl><c>, this presents the menu.
  Timing and maintenance tasks can be performed
Restart
  Command 6) measure 4ns bunch currents

## Slide 27
Support Components

## Slide 28: Support
cbpmfio library
  Management of BPM data file reading/writing
  Shared functionality between
    Server code
      C
    Analysis client code
      C
      Fortran
      Matlab (MEX wrappers)
      Python
  Open issues
    Better error checking for corrupt files is needed
    As analysis code requires access to more fields, those accessors will need to be added to the API

## Slide 29
System Development Tasks

## Slide 30: System Development Tasks
  Consolidate and refine measurement request mechanism
    Deprecate CesrBPM library
    Expand the types of control data sent via MPM
      Allow bunch pattern specification
    Best case would be to splinter off into small, single purpose request library
  Place BPM ops and related areas under version control
  Better automate the gain calibration process
    Allow polling and control of digital tune tracker
  One or more temperature fields are missing from data files

## Slide 31: System Development Tasks
  Better handling of data file splitting
    Overall, a one-measurement / one data file paradigm is simpler to manage from the analysis standpoint
  Many more…