## Sheet1
F-OPS-005: CHESS Power Outage Recovery List
Notify Chris Conolly, Katie Moring, Ernie Fontes,Zak Brown, Chris Whting, Phil Sorenson, Shijie Yang of Power Outages. Call necessary personnel as needed from the CHESS On-Call list (REF-OPS-007)
If the power stays off for >10min, start to turn the power to controllers off (MDCPs, Turbos, etc) so everything does not come on at once when power is restored. Sometimes breakers trip if MDCPs all come up at once. If that happens, turn off MDCPs and reset breaker, then power on MDCPs.
Items that may be needed to get vacuum systems back on: Jumpers, C-Cave cabinet keys, Fast VAT keys (keys found in lock box in CHESS Ops)
Notes about equipment:
To turn Terranova Ion Pump controllers on, power cycle controller. Then press and hold the HV on button for a few seconds.
To turn cold cathode gauges on: Flip HV toggle switch to OFF, power cycle controller, then flip HV toggle ON. Note: CC gauges may be found in "PRO" mode after power outage
Items in BOLD require tunnnel access
Latest Revision: 2/8/17 KMM
Turned On | System | Location | Notes
| CHESS West Chiller | Tunnel between A and B line
| A line ion pumps - Terranovas | Aline black cabinet in tunnel | 4 pumps
| A Beamstop CC gauge | Aline black cabinet in tunnel
| VBPM-A1-US | Tunnel, A-line US. Power connector under magnet. | Restarts automatically Ping server from signalcollector to verify connection.
| VBPM-A2-US | Tunnel, A-line US. Power connector under magnet. | Restarts automatically Ping server from signalcollector to verify connection.
| B line ion pumps - Terranovas | Bline black cabinet in tunnel | 4 pumps
| C line ion pumps - Terranovas | Cline black cabinet in tunnel | 7 pumps
| C line CESR-TA CC | C line black cabinet in tunnel | This signal goes to CESR Ops
| VBPM-C-DS | Tunnel, C-line DS. Power connector behind magnets | Restarts automatically Ping server from signalcollector to verify connection.
| A Cave slit ion pump - Terranova | Underneath A1 monobox | Power cycle controller, then press and hold the HV On button
| A1 F Mir ion pump - Terranova | A1 hutch, US end
| A1 focusing mirror CC | A1 hutch, US end
| A1 coffin CC gauge | A-cave (rack under A1 mono) | Controlled by PLC outside of DS A-cave door
| A1 mono box turbo pump | A-cave | Controlled by PLC outside of DS A-cave door
| VBPM-A2-DS | A-cave, Brix PC below coffin | Verify power is on. Ping server from signalcollector to verify connection
| A2 Col Mir ion pump | A-cave (rack under A2 mono) | Digitel controller. Turn power off, then on. Press and hold "Start HV" button on touch screen
| A2 Col Mir CC | A-cave (rack under A2 mono) | labeled "mir CC gauge"
| A2 coffin CC gauge | A-cave | Controlled by PLC outside of DS A-cave door
| A2 Mono box turbo pump | A-cave | Controlled by PLC outside of DS A-cave door
| AUX boxes | PLC by A-cave DS door | Press reset buttons, make sure all alarms are cleared
| VBPM-B-US | B-cave, Brix PC below coffin | Verify power is on. Ping server from signalcollector to verify connection.
| C-Coffin PLC | Above C-coffin | Answer questions on PLC as prompted
| C Cave water shutoff box | C Cave | Press reset
| C Cave turbo pump | C cave (controllers in cabinets outside of cave) | Need jumper in High Vac channel to turn turbo on
| C line upstream CC | Cabinets outside of C Cave | Labeled CC#1. Sensor is in the tunnel (not in cave)
| C Cave Cold Cathodes | Cabinets outside of C Cave | Turn on CC#2, CC#3. CC#4 can be left off
| AUX box | C Cave | Press reset button, make sure all alarms are cleared
| VBPM-C-US | C-Cave, laptop above door | Verify power is on. Ping server from signalcollector to verify connection.
| CHESS East Chiller | L0E by CHESS Helium cribs
| F cave ion pumps - Terranovas | F Cave electronics cabinet | 2 pumps - #5 and #11
| F col mir ion pump | F Cave electronics cabinet | Digitel controller, Supply 1. To turn on: Press 9, hit "select," hit "enter"
| F cave CC gauges | F Cave electronics cabinet | 2 gauges - one for col mirror, one for F1 Focusing Mirror
| Water shutoff box | F Cave, behind F1/F3 box | Press reset button, make sure all alarms are cleared
| F3 mirror water shutoff box | Outside of F cave | Press reset button, make sure all alarms are cleared
| AUX box | F Cave electronics cabinet | Press reset button, make sure all alarms are cleared
| VBPM-FH-DS | F-cave, power connector behind coffin, US | Restarts automatically. Ping server from singalcollector to verify connection
| VBPM-FW-DS | F-cave, Brix PC on top of cabinet | Verify power is on. Ping server from signalcollector to verify connection
| D line turbo | D cave | Need jumper in High Vac channel to turn turbo on
| D line US CC | D cave | Labeled "UPST CC"
| D line DS CC | D cave | Labeled "DNST CC"
| D coffin CC | D cave
| Water shutoff box | D cave
| Gate valve US and DS of coffin | D cave | Ensure they are both open, after coffin has pumped down (NOTE: Power strip on top of the rack under D-coffin may need to be power cycled to be able to get GVs to open)
| ANN Panels | D cave | Press reset buttons, make sure all alarms are cleared
| VBPM-D-US | near F3, laptop in overhead cable tray | Verify power is on. Ping server from signalcollector to verify connection
| F2 cryo system | F-cave/F2 prep area | PLC controlled - PLC in F2 prep area. LEPP Cryo & CHESS Vacuum systems both need to be working before F2 Cryo system can be restarted
| CHESS East Dump | Controls outside of F1 | Press reset button
| D-line Fast VAT | Controls outside of F1 | Valve should be open. All 5 "ready" lights should be green
| F-line Fast VAT | Controls outside of F1
| D line ion pumps - Terranovas | Dline black cabinet in tunnel | 7 pumps
| D line CESR-TA CC | Dline black cabinet in tunnel | This signal goes to CESR Ops
| E line ion pumps - Terranovas | E/F line black cabinet in tunnel | 2 pumps
| F line ion pumps - Terranovas | E/F line black cabinet in tunnel | 9 pumps
| F3 mir CC | F alcove | Duck under F line beampipes into F-alcove area. CC Controller in US black cabinet
| VBPM-FW-US | F alcove, Brix PC in cabinet under beamline | Verify power is on. Ping server from signalcollector to verify connection.
| G line chiller | G line mechanical room
| G line fan, G chem room hood, G3 exhaust | G line mechanical room | Controls immediatley to right inside mechanical room. Fan speed is normally ~50Hz
| VBPM-G-US | G Vault, brix PC in cabinet above beamline | Verify power is on. Ping server from signalcollector to verify connection
| VBPM-G-DS | G-cave, laptop in US cabletray | Verify power is on. Ping server from signalcollector to verify connection
| G line ion pumps - Terranovas | G cave rack by door | 8 pumps
| G Col Mir ion pump | G cave | Digitel MPC controller in bottom of rack just inside cave, supply 1. To turn on: Press 9, hit "select," hit "enter"
| G2/G3 FMIR pressure CC | G cave rack by door
| G1 FMIR pressure CC | G cave rack by door
| G 1st ML ion pump | G cave cab 1 | Digitel MPC, supply 1. To turn on: Press 9, hit "select," hit "enter"
| G 1st ML CC | G cave Cab 1
| G 2nd ML turbo | G cave rack by door, and cab 3 | Need jumper in High Vac channel to turn turbo on
| White beam mir box CC | G Ops | located below G2 cave control panel
| G-line VAT Valve | G Ops | Valve should be Open, All 4 "ready" lights green
| G-line Dump | G Ops | Press reset button
| G line mir box H2O Shutoff | US end of G cave | Press reset button
| G line ANN panels, AUX boxes | G Ops | Press reset buttons, make sure all alarms are cleared
| signalcollector/VBPM System | Ops, rackmount under CHESS dump panel | Power on. Open command line to ping servers. Example: ping VBPM-A1-US. Do this for: VBPM-A1-US, VBPM-A2-US, VBPM-A2-DS, VBPM-B-US, VBPM-C-US, VBPM-C-DS, VBPM-D-US, VBPM-FH-DS, VBPM-FW-DS, VBPM-FW-US, VBPM-G-US, VBPM-G-DS
| C-line VAT Valve | Ops | In CHESS mode, valve should be open, Valve "ready" light green
| CHESS West Dump | Ops | Press reset button
| ANN panels/AUX boxes | Ops | Press reset buttons, make sure all alarms have cleared
Home the undulators
| CHESS Clean Air System | PLC in 4th floor mechanical room | Need key from CESR Op to get into mechanical room. Walk up ramp, PLC is in grey cabinet on the wall on the right side as you are walking through the room. Go to "Jumpers" tab on PLC, hit "reset" then "open" for solenoid. If F2 needs clean air, will need to press "reset" and "open solenoid" from the Clean Air tab on the PLC in the F2 prep area as well.
| CESR Scoreboard video repeater | Box w/ yellow cables above CHESS West printer | If power light is flashing, unplug power cord from box then plug back in. May need to do this with same unit in CESR control room too
| Sound Meters | Near A2 and F3 | Press & Hold both the "MaxHld" and "on/off" buttons until "n" appears on the display
| Glovebox | rm 172 | Should restart on its own - verify that there are no alarms
| Chem room hood | Controls in rm 320 | Press reset button
| CHESS MDCPs | All stations, and in tunnel | Power cycle MDCPs after CHESS Private Network is up and running. When SPEC says there is a discrepency betwen the controller and what SPEC thinks the position shoudl be, go with the SPEC position.
| Misc. Equipment | G-line hallway, East roof, 141 | Check for chambers in storage under vacuum that need an ion pump turned on, or any other equipment in a temporary set-up or storage that needs attention